## Integration
6 | 1
Sample Name: |  | 1 |  |  | Injection Volume: |  | 200
Vial Number: |  | 6 |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-26 13:54 |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Height | Area | Rel.Area | Amount | Type
| min |  | µS | µS*min | % | mg/l
1 | 2.7533 | n.a. | 0.035178 | 0.0037836 | 0.013352 | n.a. | BMB
2 | 3.43 | Natrium | 1.6189 | 0.20838 | 0.73537 | 0.99449 | BMB
3 | 5.21 | Kalium | 3.2711 | 0.58843 | 2.0765 | 4.304 | BMB
4 | 9.5233 | n.a. | 56.3984 | 27.5093 | 97.0781 | n.a. | BMB
5 | 12.0333 | n.a. | 0.07962 | 0.027388 | 0.096651 | n.a. | BMB
Total: |  |  | 61.4032 | 28.3373 | 100 | 5.2985

## Calibration(Curr.Peak)
6 | 1
Sample Name: |  | 1 |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 6 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-26 13:54 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Cal.Type | Points | Corr.Coeff. | Offset | Slope | Curve
| min |  |  |  | %
1 | 2.7533 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.43 | Natrium | Lin | 4 | 99.9989 | 0 | 0.20954 | 0
3 | 5.21 | Kalium | Lin | 4 | 99.9985 | 0 | 0.13672 | 0
4 | 9.5233 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.0333 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Average: |  |  |  |  | 99.9987 | 0 | 0.17313 | 0

## Calibration(Batch)
6 | 1
Sample Name: |  | 1 |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 6 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-26 13:54 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Cal.Type | Points | Corr.Coeff. | Offset | Slope | Curve
| min |  |  |  | %
1 | 2.7533 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.43 | Natrium | Lin | 4 | 99.9989 | 0 | 0.20954 | 0
3 | 5.21 | Kalium | Lin | 4 | 99.9985 | 0 | 0.13672 | 0
4 | 9.5233 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.0333 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Average: |  |  |  |  | 99.9987 | 0 | 0.17313 | 0

## Peak Analysis
6 | 1
Sample Name: |  | 1 |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 6 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-26 13:54 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peakname | Height | Width | Type | Resol. | Asym. | Plates
| min | min | µS | min |  | (EP) | (EP) | (EP)
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Fluoride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Chloride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Nitriet | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Nitraat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Fosfaat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Sulfaat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
1 | 2.7533 | n.a. | 0.035178 | 0.16767 | BMB | 3.6349 | 0.91642 | 3590
n.a. | n.a. | Lithium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.43 | Natrium | 1.6189 | 0.192 | BMB | 7.8814 | 1.0006 | 5243
n.a. | n.a. | Ammonium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
3 | 5.21 | Kalium | 3.2711 | 0.2675 | BMB | 8.5758 | 0.97373 | 6259
n.a. | n.a. | Magnesium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 9.5233 | n.a. | 56.3984 | 0.787 | BMB | 3.7802 | 2.287 | 2613
n.a. | n.a. | Calcium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.0333 | n.a. | 0.07962 | 0.5135 | BMB | n.a. | 1.0103 | 6740
Average: |  |  | 12.2806 | 0.38553 |  | 5.9681 | 1.2376 | 4889

## SST
6 | 1
Sample Name: |  | 1 |  | Injection Volume: | 200
Vial Number: |  | 6 |  | Channel: | ECD_2
Sample Type: |  | unknown |  | Wavelength: | n.a.
Control Program: |  | Simultaneous 13 minuten |  | Bandwidth: | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  | Dilution Factor: | 1
Recording Time: |  | 2022-09-26 13:54 |  | Sample Weight: | 1
Run Time (min): |  | 13 |  | Sample Amount: | 1
System Suitability Test Results:
No. | Test Name |  | Sample Condition | Peak Condition | Test Result
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
|  | Number of executed single tests: | n.a. | Total test result: | n.a.

## Summary
Sample Name | Amount | Amount | Amount | Amount | Amount | Amount
| mg/l | mg/l | mg/l | mg/l | mg/l | mg/l
| Lithium | Natrium | Ammonium | Kalium | Magnesium | Calcium
| ECD_2 | ECD_2 | ECD_2 | ECD_2 | ECD_2 | ECD_2
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | 0.0054362
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Testsample T2t1 100x | n.a. | 0.16681 | n.a. | n.a. | n.a. | 1.1399
Testsample T2t1 100x | n.a. | 19.8143 | n.a. | 4.1088 | 0.018298 | n.a.
1 | n.a. | 0.99449 | n.a. | 4.304 | n.a. | n.a.
2 | n.a. | 19.9715 | n.a. | 4.1122 | n.a. | n.a.
3 | n.a. | 21.5366 | n.a. | 4.4739 | n.a. | n.a.
4 | n.a. | 21.589 | n.a. | 4.6623 | n.a. | n.a.
5 | n.a. | 0.32924 | n.a. | 1.6626 | n.a. | n.a.
6 | n.a. | 9.1006 | n.a. | 1.8696 | n.a. | n.a.
7 | n.a. | 8.3321 | n.a. | 1.7198 | n.a. | n.a.
8 | n.a. | 8.105 | n.a. | 1.7651 | n.a. | n.a.
9 | n.a. | 0.95186 | n.a. | 4.912 | n.a. | n.a.
10 | n.a. | 24.2497 | n.a. | 5.3551 | n.a. | n.a.
11 | n.a. | 22.9009 | n.a. | 4.9673 | n.a. | n.a.
12 | n.a. | 23.1362 | n.a. | 5.2666 | n.a. | n.a.
13 | n.a. | 0.56332 | n.a. | 2.7964 | n.a. | n.a.
14 | n.a. | 14.7496 | n.a. | 3.1969 | n.a. | n.a.
15 | n.a. | 14.2625 | n.a. | 3.0354 | n.a. | n.a.
16 | n.a. | 14.4999 | n.a. | 3.2541 | n.a. | n.a.
17 | n.a. | 0.33825 | n.a. | 1.5997 | n.a. | n.a.
18 | n.a. | 8.9683 | n.a. | 1.8905 | n.a. | n.a.
19 | n.a. | 8.4223 | n.a. | 1.7508 | n.a. | n.a.
20 | n.a. | 8.9217 | n.a. | 1.947 | n.a. | n.a.
21 | n.a. | 0.23885 | n.a. | 1.0249 | n.a. | n.a.
22 | n.a. | 6.0374 | n.a. | 1.245 | n.a. | n.a.
23 | n.a. | 5.753 | n.a. | 1.1678 | n.a. | n.a.
24 | n.a. | 6.1463 | n.a. | 1.3153 | n.a. | n.a.
25 | n.a. | 0.18196 | n.a. | 0.73593 | n.a. | n.a.
26 | n.a. | 4.4708 | n.a. | 0.91703 | n.a. | n.a.
27 | n.a. | 4.184 | n.a. | 0.85339 | n.a. | n.a.
28 | n.a. | 4.4963 | n.a. | 0.95545 | n.a. | n.a.
29 | n.a. | 0.14798 | n.a. | 0.54901 | n.a. | n.a.
30 | n.a. | 3.1088 | n.a. | 0.64822 | n.a. | n.a.
31 | n.a. | 2.9852 | n.a. | 0.61349 | n.a. | n.a.
32 | n.a. | 3.3379 | n.a. | 0.71376 | n.a. | n.a.
33 | n.a. | 0.12708 | n.a. | 0.45108 | n.a. | n.a.
34 | n.a. | 2.3278 | n.a. | 0.52401 | n.a. | n.a.
35 | n.a. | 2.3081 | n.a. | 0.50253 | n.a. | n.a.
36 | n.a. | 2.547 | n.a. | 0.57554 | n.a. | n.a.
37 | n.a. | 0.08277 | n.a. | 0.36081 | n.a. | n.a.
38 | n.a. | 1.6843 | n.a. | 0.40377 | n.a. | n.a.
39 | n.a. | 1.6643 | n.a. | 0.39411 | n.a. | 29.9139
40 | n.a. | 1.8947 | n.a. | 0.44876 | n.a. | 31.448
41 | n.a. | 0.076147 | n.a. | 0.30383 | n.a. | 25.3931
42 | n.a. | 1.2808 | n.a. | 0.33442 | n.a. | 28.6322
43 | n.a. | 1.2343 | n.a. | 0.31823 | n.a. | 26.6333
44 | n.a. | 1.4322 | n.a. | 0.35902 | n.a. | 27.5123
45 | n.a. | 0.065157 | n.a. | 0.26448 | n.a. | 26.2689
46 | n.a. | 0.94524 | n.a. | 0.27635 | n.a. | 25.1544
47 | n.a. | 0.90664 | n.a. | 0.26328 | n.a. | 23.543
48 | n.a. | 1.1234 | n.a. | 0.30738 | n.a. | 27.304
49 | n.a. | 6.7472 | n.a. | 33.8983 | 528.7972 | n.a.
50 | n.a. | 19.2516 | n.a. | 29.694 | 521.5697 | n.a.
51 | n.a. | 17.3179 | n.a. | 28.5125 | 536.545 | n.a.
52 | n.a. | 21.2454 | n.a. | 26.3109 | 553.9581 | n.a.
53 | n.a. | 411.6826 | n.a. | 87.1421 | 549.6738 | n.a.
54 | n.a. | 421.4216 | n.a. | 91.1395 | 549.2513 | n.a.
55 | n.a. | 416.1848 | n.a. | 84.3578 | 555.5313 | n.a.
Kationen 80 x | n.a. | 2.5641 | n.a. | 6.4072 | 3.371 | 6.2326
Kationen 40 x | n.a. | 5.1236 | n.a. | 12.8075 | 6.7596 | 12.4441
Kationen 20 x | n.a. | 10.2597 | 10.8632 | 25.6261 | 13.4519 | 24.8279
Kationen 10 x | n.a. | 20.6284 | 26.6438 | 51.4123 | 26.7473 | 49.9517
Anionen 40 x | n.a. | 6.2256 | n.a. | 3.1775 | n.a. | n.a.
Anionen 20 x | n.a. | 12.6127 | n.a. | 6.5373 | n.a. | n.a.
Anionen 10 x | n.a. | 24.8098 | n.a. | 12.9525 | n.a. | n.a.
Anionen 1 x | n.a. | 247.3698 | n.a. | 129.9083 | n.a. | 0.29606
Kationen 80 x | n.a. | 2.5021 | n.a. | 6.3449 | 3.3463 | 6.5241
Kationen 40 x | n.a. | 4.9724 | n.a. | 12.5716 | 6.5871 | 12.4191
Kationen 20 x | n.a. | 9.9707 | n.a. | 25.1206 | 12.8419 | 24.7094
Kationen 10 x | n.a. | 19.7557 | 25.1 | 49.6443 | 24.4515 | n.a.
Anionen 40 x | n.a. | 6.0499 | n.a. | 3.083 | 0.039018 | n.a.
Anionen 20 x | n.a. | 12.1569 | n.a. | 6.2552 | 0.078111 | n.a.
Anionen 10 x | n.a. | 24.56 | n.a. | 12.693 | 0.058893 | n.a.
Anionen 1 x | n.a. | 246.5756 | n.a. | n.a. | n.a. | n.a.
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.

## Wellness Trending
Sequence Name: | 26 september 2022
Sample Number: | 6
Sample Name: | 1
Pump Pressure Trend Plot
Background Signal Trend Plot

## Audit Trail
6 | 1 |  |  | Audit Trail
Sample Name: |  | 1 | Injection Volume: |  | 200
Vial Number: |  | 6 | Channel: |  | ECD_2
Sample Type: |  | unknown | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-26 13:54 | Sample Weight: |  | 1
Run Time (min): |  | 13 | Sample Amount: |  | 1
Day Time | Ret.Time | Command/Message
13:51:25 |  | Audit trail of sample SMP::\oude data\Ameya\26 september 2022.SEQ\4. Executed 2022/9/26 on timebase Anionen_Kationen (server PHYS9324).
13:51:25 |  | Start of sample 6 "1", using program "Simultaneous 13 minuten".
13:51:25 |  | MeasuredFlowRate = 0.3 [ml/min].
13:51:25 | 0.000 | Sampler.AcquireExclusiveAccess
13:51:25 | 0.000 | Flush Volume = 1000
13:51:25 | 0.000 | Wait FlushState
13:53:04 | 0.000 | Wait finished
13:53:04 | 0.000 | NeedleHeight = 2
13:53:04 | 0.000 | CutSegmentVolume = 10
13:53:04 | 0.000 | SyringeSpeed = 2
13:53:04 | 0.000 | CycleTime = 0
13:53:04 | 0.000 | WaitForTemperature = False
13:53:04 | 0.000 | Pump_ECD_2.Pump = On
13:53:04 | 0.000 | Pump_ECD_2.MeasuredFlowRate = 1.00
13:53:04 | 0.000 | Pump_ECD_2.Pressure.LowerLimit = 200
13:53:04 | 0.000 | Pump_ECD_2.Pressure.UpperLimit = 3000
13:53:04 | 0.000 | Pump_ECD.Pump = On
13:53:04 | 0.000 | Pump_ECD.Pressure.LowerLimit = 200
13:53:04 | 0.000 | Pump_ECD.Pressure.UpperLimit = 3000
13:53:04 | 0.000 | {Pump_ECD} Log Pump_ECD.TotalConductivity: 22.595
13:53:04 | 0.000 | {Pump_ECD} Log Pump_ECD.Pressure.Value: 2111
13:53:04 | 0.000 | {Pump_ECD_2} Log Pump_ECD_2.TotalConductivity: 1.205
13:53:04 | 0.000 | {Pump_ECD_2} Log Pump_ECD_2.Pressure.Value: 2230
13:53:04 | 0.000 | Pump_InjectValve.State = LoadPosition
13:53:04 | 0.000 | Pump_InjectValve_2.State = LoadPosition
13:53:04 | 0.000 | Wait SampleReady
13:53:05 | 0.000 | Wait finished
13:53:05 | 0.000 | Pump_ECD_2.Autozero
13:53:05 | 0.000 | Pump_ECD.Autozero
13:53:05 | 0.000 | Load
13:53:05 | 0.000 | Waiting for load response on AS50LoadState.
13:54:10 | 0.000 | Got load response
13:54:10 | 0.000 | Wait CycleTimeState
13:54:10 | 0.000 | Wait finished
13:54:10 | 0.000 | Pump_InjectValve.InjectPosition
13:54:10 | 0.000 | Pump_InjectValve_2.InjectPosition
13:54:10 | 0.000 | Inject
13:54:10 | 0.000 | Injecting from vial position 6.
13:54:10 | 0.000 | Injection Volume is 200.0 µl.
13:54:10 | 0.000 | Wait InjectState
13:54:11 | 0.000 | Wait finished
13:54:11 | 0.000 | ECD_2.AcqOn
13:54:11 | 0.000 | ECD_1.AcqOn
14:04:11 | 10.000 | Pump_InjectValve.State = LoadPosition
14:04:11 | 10.000 | Pump_InjectValve_2.State = LoadPosition
14:07:11 | 13.000 | ECD_2.AcqOff
14:07:11 | 13.000 | ECD_1.AcqOff
14:07:17 |  | End of sample "1".

## Overlay Print
Overlay of Samples from Integration View