## Integration
6 | 55 20 x diluted
Sample Name: |  | 55 20 x diluted |  |  | Injection Volume: |  | 200
Vial Number: |  | 73 |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-29 15:53 |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Height | Area | Rel.Area | Amount | Type
| min |  | µS | µS*min | % | mg/l
1 | 2.7667 | n.a. | 0.035778 | 0.0037055 | 0.023629 | n.a. | BMB
2 | 3.45 | Natrium | 33.663 | 4.3376 | 27.6595 | 20.7008 | BMB
3 | 5.2133 | Kalium | 3.1592 | 0.54842 | 3.4971 | 4.0114 | BMB
4 | 9.8367 | n.a. | 25.7083 | 10.7575 | 68.5971 | n.a. | BMb
5 | 12.06 | n.a. | 0.090028 | 0.034912 | 0.22262 | n.a. | bMB
Total: |  |  | 62.6563 | 15.6821 | 100 | 24.7122

## Calibration(Curr.Peak)
6 | 55 20 x diluted
Sample Name: |  | 55 20 x diluted |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 73 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-29 15:53 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Cal.Type | Points | Corr.Coeff. | Offset | Slope | Curve
| min |  |  |  | %
1 | 2.7667 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.45 | Natrium | Lin | 4 | 99.9989 | 0 | 0.20954 | 0
3 | 5.2133 | Kalium | Lin | 4 | 99.9985 | 0 | 0.13672 | 0
4 | 9.8367 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.06 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Average: |  |  |  |  | 99.9987 | 0 | 0.17313 | 0

## Calibration(Batch)
6 | 55 20 x diluted
Sample Name: |  | 55 20 x diluted |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 73 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-29 15:53 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peak Name | Cal.Type | Points | Corr.Coeff. | Offset | Slope | Curve
| min |  |  |  | %
1 | 2.7667 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.45 | Natrium | Lin | 4 | 99.9989 | 0 | 0.20954 | 0
3 | 5.2133 | Kalium | Lin | 4 | 99.9985 | 0 | 0.13672 | 0
4 | 9.8367 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.06 | n.a. | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
Average: |  |  |  |  | 99.9987 | 0 | 0.17313 | 0

## Peak Analysis
6 | 55 20 x diluted
Sample Name: |  | 55 20 x diluted |  |  |  | Injection Volume: |  | 200
Vial Number: |  | 73 |  |  |  | Channel: |  | ECD_2
Sample Type: |  | unknown |  |  |  | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten |  |  |  | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  |  |  | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-29 15:53 |  |  |  | Sample Weight: |  | 1
Run Time (min): |  | 13 |  |  |  | Sample Amount: |  | 1
No. | Ret.Time | Peakname | Height | Width | Type | Resol. | Asym. | Plates
| min | min | µS | min |  | (EP) | (EP) | (EP)
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. |  | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Fluoride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Chloride | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Nitriet | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Bromide | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Nitraat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Fosfaat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | Sulfaat | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
1 | 2.7667 | n.a. | 0.035778 | 0.1535 | BMB | 4.0183 | 1.2408 | 5294
n.a. | n.a. | Lithium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
2 | 3.45 | Natrium | 33.663 | 0.19183 | BMB | 7.8965 | 0.98974 | 5336
n.a. | n.a. | Ammonium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
3 | 5.2133 | Kalium | 3.1592 | 0.263 | BMB | 10.4546 | 1.0139 | 6489
n.a. | n.a. | Magnesium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
4 | 9.8367 | n.a. | 25.7083 | 0.64567 | BMb | 3.5622 | 1.5019 | 3926
n.a. | n.a. | Calcium | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
5 | 12.06 | n.a. | 0.090028 | 0.627 | bMB | n.a. | 1.0067 | 5982
Average: |  |  | 12.5313 | 0.3762 |  | 6.4829 | 1.1506 | 5405.4

## SST
6 | 55 20 x diluted
Sample Name: |  | 55 20 x diluted |  | Injection Volume: | 200
Vial Number: |  | 73 |  | Channel: | ECD_2
Sample Type: |  | unknown |  | Wavelength: | n.a.
Control Program: |  | Simultaneous 13 minuten |  | Bandwidth: | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 |  | Dilution Factor: | 1
Recording Time: |  | 2022-09-29 15:53 |  | Sample Weight: | 1
Run Time (min): |  | 13 |  | Sample Amount: | 1
System Suitability Test Results:
No. | Test Name |  | Sample Condition | Peak Condition | Test Result
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
|  | Number of executed single tests: | n.a. | Total test result: | n.a.

## Summary
Sample Name | Amount | Amount | Amount | Amount | Amount | Amount
| mg/l | mg/l | mg/l | mg/l | mg/l | mg/l
| Lithium | Natrium | Ammonium | Kalium | Magnesium | Calcium
| ECD_2 | ECD_2 | ECD_2 | ECD_2 | ECD_2 | ECD_2
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | 0.049447
Ultrapure | n.a. | n.a. | n.a. | n.a. | 0.06125 | n.a.
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.
53 20 x diluted | n.a. | 20.644 | n.a. | 4.1757 | n.a. | n.a.
54 20 x diluted | n.a. | 20.6298 | n.a. | 4.2751 | n.a. | n.a.
55 20 x diluted | n.a. | 20.7008 | n.a. | 4.0114 | n.a. | n.a.
Kationen 80 x | n.a. | 2.5641 | n.a. | 6.4072 | 3.371 | 6.2326
Kationen 40 x | n.a. | 5.1236 | n.a. | 12.8075 | 6.7596 | 12.4441
Kationen 20 x | n.a. | 10.2597 | 10.8632 | 25.6261 | 13.4519 | 24.8279
Kationen 10 x | n.a. | 20.6284 | 26.6438 | 51.4123 | 26.7473 | 49.9517
Anionen 40 x | n.a. | 6.2256 | n.a. | 3.1775 | n.a. | n.a.
Anionen 20 x | n.a. | 12.6127 | n.a. | 6.5373 | n.a. | n.a.
Anionen 10 x | n.a. | 24.8098 | n.a. | 12.9525 | n.a. | n.a.
Anionen 1 x | n.a. | 247.3698 | n.a. | 129.9083 | n.a. | 0.29606
Kationen 80 x | n.a. | 2.5021 | n.a. | 6.3449 | 3.3463 | 6.5241
Kationen 40 x | n.a. | 4.9724 | n.a. | 12.5716 | 6.5871 | 12.4191
Kationen 20 x | n.a. | 9.9707 | n.a. | 25.1206 | 12.8419 | 24.7094
Kationen 10 x | n.a. | 19.7557 | 25.1 | 49.6443 | 24.4515 | n.a.
Anionen 40 x | n.a. | 6.0499 | n.a. | 3.083 | 0.039018 | n.a.
Anionen 20 x | n.a. | 12.1569 | n.a. | 6.2552 | 0.078111 | n.a.
Anionen 10 x | n.a. | 24.56 | n.a. | 12.693 | 0.058893 | n.a.
Anionen 1 x | n.a. | 246.5756 | n.a. | n.a. | n.a. | n.a.
Ultrapure | n.a. | n.a. | n.a. | n.a. | n.a. | n.a.

## Wellness Trending
Sequence Name: | 29 september 2022
Sample Number: | 6
Sample Name: | 55 20 x diluted
Pump Pressure Trend Plot
Background Signal Trend Plot

## Audit Trail
6 | 55 20 x diluted |  |  | Audit Trail
Sample Name: |  | 55 20 x diluted | Injection Volume: |  | 200
Vial Number: |  | 73 | Channel: |  | ECD_2
Sample Type: |  | unknown | Wavelength: |  | n.a.
Control Program: |  | Simultaneous 13 minuten | Bandwidth: |  | n.a.
Quantif. Method: |  | Barbara 19 oktober 2019 | Dilution Factor: |  | 1
Recording Time: |  | 2022-09-29 15:53 | Sample Weight: |  | 1
Run Time (min): |  | 13 | Sample Amount: |  | 1
Day Time | Ret.Time | Command/Message
15:49:49 |  | Audit trail of sample SMP::\oude data\Ameya\29 september 2022\29 september 2022.SEQ\17079. Executed 2022/9/29 on timebase Anionen_Kationen (server PHYS9324).
15:49:49 |  | Start of sample 6 "55 20 x diluted", using program "Simultaneous 13 minuten".
15:49:49 |  | MeasuredFlowRate = 0.3 [ml/min].
15:49:49 | 0.000 | Sampler.AcquireExclusiveAccess
15:49:49 | 0.000 | Flush Volume = 1000
15:49:49 | 0.000 | Wait FlushState
15:51:27 | 0.000 | Wait finished
15:51:27 | 0.000 | NeedleHeight = 2
15:51:27 | 0.000 | CutSegmentVolume = 10
15:51:27 | 0.000 | SyringeSpeed = 2
15:51:27 | 0.000 | CycleTime = 0
15:51:27 | 0.000 | WaitForTemperature = False
15:51:27 | 0.000 | Pump_ECD_2.Pump = On
15:51:27 | 0.000 | Pump_ECD_2.MeasuredFlowRate = 1.00
15:51:27 | 0.000 | Pump_ECD_2.Pressure.LowerLimit = 200
15:51:27 | 0.000 | Pump_ECD_2.Pressure.UpperLimit = 3000
15:51:27 | 0.000 | Pump_ECD.Pump = On
15:51:27 | 0.000 | Pump_ECD.Pressure.LowerLimit = 200
15:51:27 | 0.000 | Pump_ECD.Pressure.UpperLimit = 3000
15:51:27 | 0.000 | {Pump_ECD} Log Pump_ECD.TotalConductivity: 22.594
15:51:27 | 0.000 | {Pump_ECD} Log Pump_ECD.Pressure.Value: 2095
15:51:27 | 0.000 | {Pump_ECD_2} Log Pump_ECD_2.TotalConductivity: 1.109
15:51:27 | 0.000 | {Pump_ECD_2} Log Pump_ECD_2.Pressure.Value: 2207
15:51:27 | 0.000 | Pump_InjectValve.State = LoadPosition
15:51:27 | 0.000 | Pump_InjectValve_2.State = LoadPosition
15:51:27 | 0.000 | Wait SampleReady
15:51:29 | 0.000 | Wait finished
15:51:29 | 0.000 | Pump_ECD_2.Autozero
15:51:29 | 0.000 | Pump_ECD.Autozero
15:51:29 | 0.000 | Load
15:51:29 | 0.000 | Waiting for load response on AS50LoadState.
15:52:33 | 0.000 | Got load response
15:52:33 | 0.000 | Wait CycleTimeState
15:52:33 | 0.000 | Wait finished
15:52:33 | 0.000 | Pump_InjectValve.InjectPosition
15:52:33 | 0.000 | Pump_InjectValve_2.InjectPosition
15:52:33 | 0.000 | Inject
15:52:33 | 0.000 | Injecting from vial position 73.
15:52:33 | 0.000 | Injection Volume is 200.0 µl.
15:52:33 | 0.000 | Wait InjectState
15:52:34 | 0.000 | Wait finished
15:52:34 | 0.000 | ECD_2.AcqOn
15:52:34 | 0.000 | ECD_1.AcqOn
16:02:34 | 10.000 | Pump_InjectValve.State = LoadPosition
16:02:34 | 10.000 | Pump_InjectValve_2.State = LoadPosition
16:05:34 | 13.000 | ECD_2.AcqOff
16:05:34 | 13.000 | ECD_1.AcqOff
16:05:41 |  | End of sample "55 20 x diluted".

## Overlay Print
Overlay of Samples from Integration View